Request For Quotes for Micro-Focusing Silicon Mirrors

for

the MHATT-CAT 7BM beamline

at the

Advanced Photon Source

July 14, 2004

Eric Dufresne and Don Walko

MHATT-CAT, APS Sector 7,

Argonne National Lab Bldg 432D

9700 South Cass Ave

Argonne IL 60439

630-252-0274

Please email quote to both: dufresne@umich.edu and d-walko@anl.gov.

Copied from a version sent by

Peter J. Eng on September 10, 2002

Consortium for Advanced Radiation Sources

The University of Chicago

5640 South Ellis Avenue

Chicago, Illinois 60637

(773) 702-2064

eng@cars.uchicago.edu

# Introduction

We are requesting two silicon mirrors: two 200mm long mirrors -- the exact mirror dimensions are given in section 4.  All mirror shall be fabricated according to general procedure outlined below, the final specification of roughness and slope error for each type is given in section 3. All these specifications are the same as those of the mirrors fabricated for Peter Eng.

# General Fabrication Procedure:

- The mirrors shall be fabricated from a boule of single crystal CZ Si with no slip lines across the length and supplied by the Vender.

- Mirror blanks are to be ground to the proper trapezoidal shape and thickness adding extra thickness so that the final polished mirror thickness is with in 5% of specified thickness.

- Bevel all edges.

- Acid etched blanks removing approximately 0.003” from all surfaces.

- Lap and polish back side of mirror to a commercial finish.

- Lap and polish optical surface to the specification in section 3.

# Specifications

Each mirror type is identified by its length L, focal length F and thickness T.  For each mirror a maximum slope error and roughness is given, along with minimum convex and concave radius of curvature.

| Qty | Mirror Type | Maximum Slope Error1 [rad] | Maximum RMS Roughness [Å] | Minimum Convex Radius [km] | Minimum Concave Radius [km] |
| --- | --- | --- | --- | --- | --- |
| 1 | L = 200 mm F = 190 mm T = 6.25 mm Fig. 3 | 1.5 | 1.0 | 0.5 | 1.5 |
| 1 | L = 200 mm F = 390 mm T = 7.0 mm Fig. 4 | 1.5 | 1.0 | 0.5 | 2.0 |

(1) With first order curvature (power) removed.

# Mirror Shapes

The figures below show dimensions of the trapezoidal mirrors and the last figure defines the concave and convex radius.  The optical surface over which the specification given in section 3 must be held is shown as the hatched area.

Fig. 1  Horizontal focusing mirror 100mm long.

Fig. 2  Vertical focusing mirror 100 mm long.

Fig. 3  Horizontal focusing mirror 200 mm long.

Fig. 4  Vertical focusing mirror 200 mm long.

Fig. 5  Definition of convex and concave radius.