Description of the measured component:

Flat trapezoid Si mirror:

  - Surface size: 100 mm x 42 mm x 20 mm

  - Coating: Pt

Performed measurements:

- Surface roughness

- Surface slope

Instruments used for measuring:

An “X” indicates the instrument(s) used.

| Surface roughness: |  | WYKO-TOPO-2D, |  | 1.5x | X | 5x |  | 40x |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | WYKO-TOPO-3D, |  | 1.5x |  | 5x |  | 40x |
| Surface figure: |  | WYKO-6000 |  |  |  |  |  |  |
| Surface slope: | X | LTP |  |  |  |  |  |  |

###### Surface roughness measurement with the TOPO 2D, objective 5X

- The separation of adjacent measured spots was about 10 mm.

- Table 1 summarizes the roughness measurement results of the mirror.

- Figure 1 shows an example of the surface roughness profiles.

Table 1. Results of the roughness measurements. (Profile length: 2.66 mm)

| Meas. (#) | RMS (nm) | P-V (nm) | Radius (m) | Figure (#) |
| --- | --- | --- | --- | --- |
| 1 | 0.151 | 1.50 | 1254 | 1 |
| 2 | 0.145 | 1.06 | 879 |  |
| 3 | 0.109 | 0.78 | 1013 |  |
| 4 | 0.208 | 1.85 | 906 |  |
| 5 | 0.121 | 1.40 | 928 |  |
| Mean | 0.147 | 1.32 |  |  |

### Figure 1. An example of the roughness profiles (measurement 1).

### Surface slope profile measurement with the LTP

- The mirror was set face-up and scanned over a length of 96 mm along the central line (see Figure 2).

- Table 2 summarizes the results of the LTP measurements. Residual slope data were obtained by removing the best-fit cylinder from the LTP raw profiles.

- Figures 3 and 4 show the residual slope and height profiles with different lengths.

72 mm

Scan direction

Figure 2. Scan direction.

### Table 2. Results of the surface slope measurements.

| Profile length (mm) | Residual slope |  | Residual height |  | Radius (km) | Figure (#) |
| --- | --- | --- | --- | --- | --- | --- |
|  | RMS (rad) | P-V (rad) | RMS (m) | P-V (m) |  |  |
| 96 | 2.66 | 14.15 | 0.05 | 0.08 | 2.42 | 3 |
| 72 | 1.25 | 5.69 | 0.02 | 0.03 | 1.98 | 4 |

Figure 3. Residual slope and height profiles of the mirror. (Length: 96 mm)

Figure 4. Residual slope and height profiles of the mirror. (Length: 72 mm)